## Instructions
The instructions, Breakpoints in Use template, and examples (ie, "Demo Data") provided herein are suggestions for documenting breakpoints in use.
Each laboratory should edit the form or use terminology that differs from that suggested here, as appropriate.
Suggested procedure for completing “Breakpoints in Use” Form:
1. Arrange a meeting with an appropriate IT staff member in your facility to confirm where (besides in an automated AST instrument software) breakpoints may be currently stored and applied at your institution. At many institutions, breakpoints are often also built into the LIS and/or EHR.
2. If using a commercial AST system, ask your system’s AST technical representative for instructions on how to obtain a list of breakpoints being applied on your system or refer to your instrument/software user manual.
3. For drugs currently tested within your lab, compare breakpoints being used by your lab to those in the current edition of CLSI's document M100. Flag the breakpoints being used in your laboraory that differ from the current breakpoints (found in CLSI document M100).
4. Cross-check breakpoints that are flagged in #3 with breakpoints provided on the FDA susceptibility test interpretive category (STIC [BPs]) website listed below to see if CLSI breakpoints = FDA breakpoints.
https://www.fda.gov/drugs/development-resources/antibacterial-susceptibility-test-interpretive-criteria
a. If CLSI breakpoints = FDA breakpoints but are different from those in use in your laboratory:
Develop a plan for implementing updated breakpoints. This might involve meeting with your antimicrobial stewardship program (ASP) team to prioritize updates (if multiple breakpoint updates for different drugs are required) and review reporting needs for the drug(s).
b. If CLSI breakpoints ≠ FDA breakpoints: Meet with your ASP to discuss which breakpoints are appropriate for your facility.
5. Develop a plan (including timeline) to update any breakpoints in use that do not coincide with CLSI M100 (current version) and/or FDA STIC (breakpoints).
Notes About Variables Suggested in Columns in the "Breakpoints in Use Template" Sheet
Column: Location of Breakpoints
● Automated instruments likely house breakpoints that will automatically interpret MIC results.
● Disk diffusion measurements may be interpreted manually prior to entry into an LIS; in this case, source of breakpoints is likely referenced in the SOP.
● Disk diffusion measurements may be interpreted automatically in the LIS or EHR.
● Interpretive results for some drugs generated with an instrument may be overridden and interpreted manually prior to entry into an LIS; in this case, source of BPs is likely referenced in the SOP.
Column: Breakpoint Matches Current M100 as of Date of Lab Review?
● The current edition of CLSI document M100 is the most recent edition as listed on CLSI's website (clsi.org).
● Breakpoints listed match those published in the current edition of CLSI document M100.
Column: Breakpoint Matches FDA STIC as of Date of Lab Review?
● Current FDA Breakpoints are found here:
https://www.fda.gov/drugs/development-resources/antibacterial-susceptibility-test-interpretive-criteria
● Breakpoints listed match those published on the FDA STIC website on the Date of Lab Review.
Column: Date Breakpoint Implemented in Lab
● Date when the breakpoint was incorporated into the laboratory procedures; if unknown, it is acceptable to list "pre-2021," the year in which recording breakpoints became a CAP requirement.
Column: Date of Lab Review
● The date on which the laboratory either reviewed and/or affirmed the breakpoints in use as listed on this spreadsheet.
Abbreviations
ASP, antimicrobial stewardship program
EHR, electronic health record where final laboratory reports are posted
LIS, laboratory information system
SOP, standard operating procedure (laboratory procedure)
STIC, susceptibility test interpretive criteria (FDA terminology for breakpoints)
ZD, zone diameter
CLSI Version 1.0. This was last updated on 15 May 2023 and has been approved by CLSI’s Outreach Working Group.
This is Part A of the 2023 Breakpoint Implementation Toolkit. To access the entire BIT Toolkit, visit https://clsi.org/bit-toolkit/.

## Breakpoints in Use template
Antimicrobial Agent | Organism/Group | Test System | Interpretive Categories and MIC Breakpoints (μg/mL) or Zone Diameter Breakpoints (nearest whole mm) |  |  |  | Location of Breakpoint (instrument/ LIS/SOP/EHR) | Brekapoint matches current M100 as of lab review date? | Breakpoint matches FDA STIC as of lab review date? | Date Breakpoints implemented in lab | Date of lab review | Comments/Action Plan
|  |  | Susceptible, MIC ≤ or ZD ≥ | Susceptible Dose-Dependent | Intermediate | Resistant, MIC ≥ or ZD ≤

## Demo Data
Antimicrobial Agent | Organism/Group | Test System | Interpretive Categories and MIC Breakpoints (μg/mL) or Zone Diameter Breakpoints (mm) |  |  |  | Location of Breakpoint (instrument/ LIS/SOP/EHR) | Breakpoint matches current M100 as of lab review date? | Breakpoint matches FDA STIC as of lab review date? | Date Breakpoints implemented in lab | Date of lab review | Comments/Action Plan
|  |  | Susceptible, MIC ≤ or ZD ≥ | Susceptible Dose-Dependent | Intermediate | Resistant, MIC ≥ or ZD ≤
Cefepime | Enterobacterales | Commercial automated device | 2 | 4-8 | n/a | 16 | LIS | Yes | No | Pre-2021 | 2022-05-12 | Instrument follows FDA STIC, categorizing 4-8 μg/mL as Intermediate. LIS converts intermediate to CLSI SDD categorization.
Cefepime | Enterobacterales | Disk diffusion | 25 | 19-24 | n/a | 18 | EHR | Yes | No | Pre-2021 | 2022-05-12
Cefepime | P. aeruginosa | Commercial automated device | 8 | n/a | 16 | 32 | LIS | Yes | No | Pre-2021 | 2022-05-12
Cefepime | P. aeruginosa | Disk diffusion | 18 | n/a | 15-17 | 14 | EHR | Yes | No | Pre-2021 | 2022-05-12
Ceftazidime | Enterobacterales | Commercial automated device | 8 | n/a | 16 | 32 | LIS | No | No | Pre-2021 | 2022-05-12 | Obsolete breakpoint, must update; see validation plan. Antimicrobial not routinely reported. Currently test by disk diffusion if requested.
Ceftazidime | Enterobacterales | Disk diffusion | 21 | n/a | 18-20 | 17 | EHR | Yes | Yes | Pre-2021 | 2022-05-12
Ceftazidime | P. aeruginosa | Commercial automated device | 8 | n/a | 16 | 32 | LIS | Yes | No | Pre-2021 | 2022-05-12
Ceftazidime | P. aeruginosa | Disk diffusion | 18 | n/a | 15/17 | 14 | EHR | Yes | No | Pre-2021 | 2022-05-12
Ciprofloxacin | Enterobacterales | Commercial automated device | 0.25 | n/a | 0.5 | 1 | LIS | Yes | Yes | Pre-2021 | 2022-05-12 | In-use for Enterobacterales except Salmonella. Panel concentrations are not low enough for Salmonella.
Ciprofloxacin | Enterobacterales | Disk diffusion | 26 | n/a | 22-25 | 21 | EHR | Yes | Yes | Pre-2021 | 2022-05-12
Ciprofloxacin | Salmonella spp. | Gradient strip | 0.06 | n/a | 0.12-0.5 | 1 | SOP | Yes | Yes | Pre-2021 | 2022-05-12
Ciprofloxacin | P. aeruginosa | Commercial automated device | 0.5 | n/a | 1 | 2 | LIS | Yes | Yes | Pre-2021 | 2022-05-12
Ciprofloxacin (direct blood culture 8-10hr) | P. aeruginosa | Disk diffusion | 23 | n/a | 18-22 | 17 | EHR | Yes | No | 2022-05-03 | 2022-05-12
Ciprofloxacin (direct blood culture 16-18hr) | P. aeruginosa | Disk diffusion | 25 | n/a | 19-24 | 18 | EHR | Yes | No | 2022-05-03 | 2022-05-12
Daptomycin | Enterococcus spp. | Commercial automated device | 4 | n/a | n/a | n/a | LIS | No | No | Pre-2021 | 2022-05-12 | Obsolete breakpoint, must update; see validation plan. Instrument came with former FDA breakpoint <=4 S reportable only for E. faecalis, and an in-house validation was performed in 2016 to expand reporting to all Enterococcus spp. using former CLSI breakpoint of <=4 S. Must now validate to implement current separate CLSI breakpoints for E. faecalis, breakpoints for E. faecium, and breakpoints for other Enterococcus spp.
Daptomycin | Enterococcus spp. | Gradient strip | 4 | n/a | n/a | n/a | SOP | No | No | Pre-2021 | 2022-05-12 | Obsolete breakpoint, must update; see validation plan. Instrument came with former FDA breakpoint <=4 S reportable only for E. faecalis, and an in-house validation was performed in 2016 to expand reporting to all Enterococcus spp. using former CLSI breakpoint of <=4 S. Must now validate to implement current separate CLSI breakpoints for E. faecalis, breakpoints for E. faecium, and breakpoints for other Enterococcus spp
Action needed